## Slide 1: Setting up BoxCar-Like methods on a Fusion
Ben Orsburn
www.LCMSMethods.org

## Slide 2: Disclaimers
1) Use at your own risk
2) No guarantees of success
3) Topical use only
4) All technologies described are the properties of other people and the companies, universities, institutes that developed them. All their trademarks apply. I own 2 dogs I got new, 2 dogs I got used with a load of miles on them, and a cat I found in the woods of Pennsyltucky that was all messed up and lives upstairs for ever now, I guess. I don’t own one single patent on anything. Please only use this and anything found on www.LCMSmethods.org for the good of mankind, please.

## Slide 3: Many BoxCar-like methods can be set up on the Tribrids. I group them as 3 things
1) BoxCar method that most resembles the original study (BoxFahrt)
2) BoxCar Assisted MS2 Fragmentation (BAMF) – the use of BoxCars to actively select the ions used for MS/MS. At present, I believe this can only be performed on the Fusion and Lumos devices
3) BoxCarDIA – BoxCar in conjunction with Data Independent Acquisition – this is a test method that does appear to work, but I can’t find anything that can process the data.

## Slide 4: How to set up any BoxCar like method on a Tribrid (Fusion Tune 3.0 shown)
Tribrid instruments can have multiple “Experiments”. Each MS1 or MS2 acquisition type requires it’s own “Experiment”
In the Example below (BAMF), I have 3 Experiments. Each Experiment can be configured independently.

## Slide 5: BoxCar is essentially Multiplex Targeted-SIM (msx-TSIM)
-msx allows isolation and collection prior to a scan
-factory instrument parameters limit you to 10 msx per scan
-BoxCar in MaxQuant.Live does not have this limitation. In the original BoxCar study 12 windows are utilized
Each filling can have it’s own target!!

## Slide 6: BoxFahrt methods
BoxFahrt methods come close to replicating the original BoxCar method.
This method requires 4 separate “Experiments”
3 BoxCars are used to acquire more democratic precursor distributions.
A normal MS1 with ddMS2 triggered off of the MS1 is the final 4th experiment

## Slide 7: BoxCar Assisted MS/MS Fragmentation (BAMF)
BAMF utilizes only 3 experiments.
Ions selected for fragmentation are chosen directly from the BoxCar scans rather than the MS1
As such, ions are selected for fragmentation democratically, allowing MS/MS fragmentation and sequencing of ions that can not be seen in a “regular” full MS1 scan

## Slide 8: BoxCarDIA
This is an experimental method that does appear to work, however I have had no luck as of ASMS 2019 finding anything that can process the data. Please contact me if you would like to have a RAW file.
The method utilizes 6 experiments alternating between 3 BoxCar (msx-TSIM) scans and 3 small window DIA scans.
Using the ion trap in parallel with the Orbitrap effectively allows DIAs to be acquired with no increase in cycle time