Nucleic Acid Extraction Using the MagMAX™ CORE Nucleic Acid Purification Kit on a Magnetic Particle Processor Worksheet

This worksheet to be used in conjunction with Nucleic Acid Extraction Using the MagMAX™ CORE Nucleic Acid Purification Kit on a Magnetic Particle Processor Protocol

Date: Click or tap to enter a date.___Operator:      ____________________________________________

Project/ Accession Number:

Step 1.  Extraction Reagent Preparation

Fill in Lot # and Initials.

Note:  Enter the number of samples in the highlighted well under “Rxns”.

Volumes needed will automatically be calculated with an additional 10% overage.

Step 2. Extraction Plate Preparation/ Procedure

- Reagent Preparation

| Prepare extraction reagents in a hood in the clean room according to the chart above. Note: Xeno™ RNA Control should be added to the Lysis/Binding Solution in the extraction room. | ☐ |
| --- | --- |

- Plate Preparation:

| Microplate 1. Wash Solution 1 |  |  |
| --- | --- | --- |
| 1. | In clean room, add 500µl of Wash Solution 1 to a deep well microplate. | ☐ |
| Microplates 2. Wash Solution 2 |  |  |
| 2. | In clean room, add 500µl of Wash Solution 2 to a deep well microplate. | ☐ |
| Microplate 3. Elution |  |  |
| 3. | In clean room, add 90µl of Elution Buffer to a standard (or deep well) microplate. | ☐ |
| Microplate 4. Beads/ PK Mix, Sample, Lysis, and Xeno™ RNA |  |  |
| 4. | In clean room, add 30µl Bead/PK mix to a deep well microplate. | ☐ |
| 5. | In extraction room, add 200µl of each prepared sample (For whole blood samples: add 100µl water and 100µl blood sample) to a deep well microplate. Cover and place on plate shaker for ~2 minutes at moderate speed. | ☐ |
| 6. | In extraction room, add 700µl of Lysis/Binding Solution to each sample. Immediately proceed to process samples on the magnetic particle processor. | ☐ |

- Load Magnetic Particle Processor (Note: Plates are loaded in reverse order of procedure.)

| Position 5 | Load empty standard (or deep well) microplate and deep well rod cover |  | ☐ |
| --- | --- | --- | --- |
| Position 4 | Load standard (or deep well) Microplate 6 (Elution) |  | ☐ |
| Position 3 | Load deep well microplate 4 (Wash Solution 2) |  | ☐ |
| Position 2 | Load deep well microplate 2 (Wash Solution 1) |  | ☐ |
| Position 1 | Load deep well microplate 1 (Beads/PK, Sample, Lysis, and Xeno™ RNA) |  | ☐ |

- Start Run

| Start run on the magnetic particle processor. Protocol: MagMAX_CORE_Flex.bdz | ☐ |
| --- | --- |
| When run is completed, remove Elution Plate from magnetic particle processor. Eluted nucleic acid can be transferred to an alternate storage container if desired. Use immediately or store the nucleic acid at -60 or cooler until use. | ☐ |
| Complete worksheet and enter in appropriate notebook/data collection log. | ☐ |

- Deviations (Any deviations must be pre-approved, noted, and initialed in the box below.)

Signature: ______________________________________________ Date: ____________________________